## Slide 1: Anesthetic Monitoring
OBJECTIVES
At the end of the lecture you will be able to know the basics of anesthetic monitoring as follows:
Definition
Where, when, what to monitor
The rules and regulations that govern modern monitoring
The basic monitors and the advanced monitors
Arterial Oxygen Saturation- SpO2
Expired CO2 - ETCO2
Awareness under anesthesia
Means to monitor the wakeful state of the brain
Other somatosensory and motor monitoring
The neuromuscular junction relaxation monitoring
Brief introduction about invasive hemodynamic monitoring and oxygenation of the brain

## Slide 2: Anesthetic Monitoring
Definition
What is Monitoring ?

## Slide 3: Anesthetic Monitoring
Definition
What is Monitoring ?
observe and check the progress or quality of (something) over a period of time.
keep under systematic review.

## Slide 4: Anesthetic Monitoring
What do you Monitor in a patient?

## Slide 5: Anesthetic Monitoring
What do you Monitor in a patient?
Vitals
Color/skin
Wakefulness

## Slide 6: Anesthetic Monitoring
How and by which means do you Monitor in a patient?

## Slide 7: Anesthetic Monitoring
How and by which means do you Monitor in a patient?
Physical exam
Equipments ( advances in technology)

## Slide 8: Anesthetic Monitoring
What are the Standards to follow for monitoring a patient
Responsibilities?

## Slide 9: Anesthetic Monitoring
What determines the Standards of Care for monitoring a patient
Basic monitoring
Advanced monitoring

## Slide 10: Anesthetic Monitoring
What determines the Standards of Care for monitoring a patient
Patient/ illness
Equipments/ technology
Rules/ legislation

## Slide 11: Anesthetic Monitoring
What determines the Standards of Care for monitoring a patient
Anesthesia type?
General
Regional/neuraxial
Monitored Anesthesia Care/Sedation

## Slide 12: Anesthetic Monitoring
What is Anesthesia?
Hypnosis
Analgesia
Paralysis

## Slide 13: Anesthetic Monitoring
What would happen for the body during anesthesia?
Neurodepression/respiratory
Cardiodepression/BP CO
Vasodilation
Low BP affects perfusion to vital organs
Low Oxygen affects metabolism of organs

## Slide 14: Anesthetic Monitoring
Standards for Anesthetic Monitoring

## Slide 15: Anesthetic Monitoring
Standards for Anesthetic Monitoring
These standards
apply to all anesthesia care although, in emergency circumstances, appropriate life support measures take precedence.
may be exceeded at any time based on the judgment of the responsible anesthesiologist.
They are intended to encourage quality patient care, but observing them cannot guarantee any specific patient outcome.
They are subject to revision from time to time, as warranted by the evolution of technology and practice.
They apply to all general anesthetics, regional anesthetics and monitored anesthesia care.

## Slide 16: Anesthetic Monitoring
Standards for Anesthetic Monitoring
This set of standards addresses only the issue of basic anesthetic monitoring, which is one component of anesthesia care.
In certain rare or unusual circumstances,
1) some of these methods of monitoring may be clinically impractical, and
2) appropriate use of the described monitoring methods may fail to detect untoward clinical developments.

## Slide 17: Anesthetic Monitoring
Standards for Anesthetic Monitoring
Brief interruptions of continual monitoring may be unavoidable.
Note that “continual” is defined as “repeated regularly and frequently in steady rapid succession” whereas “continuous” means “prolonged without any interruption at any time.”

## Slide 18: Anesthetic Monitoring
Standards for Anesthetic Monitoring
Standard I
Qualified anesthesia personnel shall be present in the room throughout the conduct of all general anesthetics, regional anesthetics and monitored anesthesia care.
Due to the rapidity of occurrence of physiologic derangement during surgical interference

## Slide 19: Anesthetic Monitoring
Standards for Anesthetic Monitoring
Standard I
In the event there is a direct known hazard, e.g., radiation, to the anesthesia personnel which might require intermittent remote observation of the patient, some provision for monitoring the patient must be made.

## Slide 20: Anesthetic Monitoring
Standards for Anesthetic Monitoring
Standard I
In the event that an emergency requires the temporary absence of the person primarily responsible for the anesthetic, the best judgment of the anesthesiologist will be exercised in comparing the emergency with the anesthetized patient’s condition and in the selection of the person left responsible for the anesthetic during the temporary absence.

## Slide 21: Anesthetic Monitoring
Standards for Anesthetic Monitoring
Standard II
During all anesthetics, the patient’s oxygenation, ventilation, circulation and temperature shall be continually evaluated

## Slide 22: Anesthetic Monitoring
Standard II
Frequency of mandatory monitoring varies between each category, but never exceeds five minutes.
Standards for Anesthetic Monitoring

## Slide 23: Anesthetic Monitoring
Standard II
Frequency of mandatory monitoring varies between each category, but never exceeds five minutes.
If not used, a reason should be recorded on the patient record.
Standards for Anesthetic Monitoring

## Slide 24: Anesthetic Monitoring
Standard II
The following are all specifically mandated.
Standards for Anesthetic Monitoring

## Slide 25: Anesthetic Monitoring
Standard II
The following are all specifically mandated.
Oxygen analyzer with a low inspired concentration limit alarm during general anesthesia
Standards for Anesthetic Monitoring

## Slide 26: Anesthetic Monitoring
Standard II
The following are all specifically mandated.
Oxygen analyzer with a low inspired concentration limit alarm during general anesthesia
Most modern anesthesia machines monitor both inspired and expired concentrations of O2
This is essential during anesthesia because it is possible to deliver a hypoxic gas mixture when mixing O2, air, nitrous oxide, and/or volatile anesthetic agents.
Standards for Anesthetic Monitoring

## Slide 27: Anesthetic Monitoring
Standard II
The following are all specifically mandated.
Oxygen analyzer with a low inspired concentration limit alarm during general anesthesia
Quantitative assessment of blood oxygenation
Standards for Anesthetic Monitoring

## Slide 28: Anesthetic Monitoring
Standard II
The following are all specifically mandated.
Oxygen analyzer with a low inspired concentration limit alarm during general anesthesia
Quantitative assessment of blood oxygenation
Pulse Oximetry
Provides quantitative analysis of the patient's saturation of hemoglobin with O2.
Standards for Anesthetic Monitoring

## Slide 29: Anesthetic Monitoring
Standard II
The following are all specifically mandated.
Oxygen analyzer with a low inspired concentration limit alarm during general anesthesia
Quantitative assessment of blood oxygenation
Ensuring adequate ventilation during all anesthetic care including verification of expired oxygen (when possible), quantitative measurement of tidal volume, and capnography in all general anesthetics.
Standards for Anesthetic Monitoring

## Slide 30: Anesthetic Monitoring
Standard II
The following are all specifically mandated.
Qualitative evaluation of ventilation is required during all other care.
Standards for Anesthetic Monitoring

## Slide 31: Anesthetic Monitoring
Standard II
The following are all specifically mandated.
Qualitative evaluation of ventilation is required during all other care.
  Inspired and expired CO2 should be monitored.
  Expired CO2 is frequently displayed through capnography with a displayed value correlating to the peak expired CO2 of each breath
Standards for Anesthetic Monitoring

## Slide 32: Anesthetic Monitoring
Standard II
The following are all specifically mandated.
Qualitative evaluation of ventilation is required during all other care.
Ensure correct placement of endotracheal tube or laryngeal mask airway via expired carbon dioxide (CO2).
Standards for Anesthetic Monitoring

## Slide 33: Anesthetic Monitoring
Standard II
The following are all specifically mandated.
Qualitative evaluation of ventilation is required during all other care.
Ensure correct placement of endotracheal tube or laryngeal mask airway via expired carbon dioxide (CO2).
Alarms for disconnects when a mechanical ventilator is used
Standards for Anesthetic Monitoring

## Slide 34: Anesthetic Monitoring
Standard II
The following are all specifically mandated.
Continuous display of ECG
Standards for Anesthetic Monitoring

## Slide 35: Anesthetic Monitoring
Standard II
The following are all specifically mandated.
Continuous display of ECG
Determination of arterial BP and heart rate at least every 5 minutes.
Standards for Anesthetic Monitoring

## Slide 36: Anesthetic Monitoring
Standard II
The following are all specifically mandated.
Continuous display of ECG
Determination of arterial BP and heart rate at least every 5 minutes.
Adequacy of circulation is to be determined by quality of pulse either electronically, through palpation, or auscultation
Standards for Anesthetic Monitoring

## Slide 37: Anesthetic Monitoring
Standard II
The following are all specifically mandated.
Continuous display of ECG
Determination of arterial BP and heart rate at least every 5 minutes.
Adequacy of circulation is to be determined by quality of pulse either electronically, through palpation, or auscultation
The means to determine temperature must be available and should be employed when changes in temperature are anticipated or intended.
Standards for Anesthetic Monitoring

## Slide 38: Anesthetic Monitoring
Multiple expired gas analysis
Allows determination of the percent inspired and expired of the volatile agents and nitrous oxide.
This allows the ability to better determine the delivery of an adequate anesthetic without over or under dose.
Standards for Anesthetic Monitoring

## Slide 39: Anesthetic Monitoring
ECG
The minimum of three leads is to be used, although five leads are used for most adults.
Consideration must be taken for the surgical field and patient positioning.
  Lead placement is commonly altered for cases involving the chest, shoulders, back, and neck.
Standards for Anesthetic Monitoring

## Slide 40: Anesthetic Monitoring
ECG
Five Lead ECG
    Includes the right arm (RA), left arm (LA), right leg (RL), left leg (LL), and V.
    The five lead arrangement can be used to display I, II, III, aVR, aVL, aVF, and/or V
Three lead ECG
    Includes the RA, LA, and LL leads and can be used to display leads I, II,and/or Ill
Standards for Anesthetic Monitoring

## Slide 41: Anesthetic Monitoring
Standards for Anesthetic Monitoring

## Slide 42: Anesthetic Monitoring
A three lead ECG can be modified to display V5 by moving the LA lead to the V5 position in the fifth intercostal space at the anterior axillary line
Standards for Anesthetic Monitoring
ECG

## Slide 43: Anesthetic Monitoring
The most commonly monitored leads are II and V5
II is best used to monitor rhythm because it provides the best visibility of the P wave
V5 monitors for anterior and lateral ischemic events
Standards for Anesthetic Monitoring
ECG

## Slide 44: Anesthetic Monitoring
If an arrhythmia or ischemic event appears to be present, the ability to viewing all leads simultaneously may be helpful for diagnostic purposes.
Standards for Anesthetic Monitoring
ECG

## Slide 45: Anesthetic Monitoring
Arterial blood pressure (BP)
  BP can be monitored non-invasively or invasively.
  Non-invasive methods include oscillometric cuff , and rarely palpation, ausculatation, Doppler probe.
Standards for Anesthetic Monitoring

## Slide 46: Anesthetic Monitoring
Arterial blood pressure (BP)
Automatic oscillometric
  The cuff is able to sense oscillations in cuff pressure which correlate with arterial pulsation.
  Placement
    Each cuff is labeled with an arrow pointing to where arterial pulsation is felt best.
    The cuff is then placed on the arm over the brachial artery, forearm over the radial artery, or thigh/calf over the popliteal artery.
  Patient positioning
    When monitoring non-invasive pressure, consideration must be taken of patient position.
Standards for Anesthetic Monitoring

## Slide 47: Anesthetic Monitoring
7) Arterial blood pressure (BP)
Standards for Anesthetic Monitoring

## Slide 48: Anesthetic Monitoring
Arterial blood pressure (BP)
Standards for Anesthetic Monitoring

## Slide 49: Anesthetic Monitoring
  Invasive pressure monitoring
  Arterial : allows for continuous beat to beat monitoring of arterial blood pressure displayed as a waveform and provides access for arterial sampling
Modalities for Anesthetic Monitoring

## Slide 50: Anesthetic Monitoring
Temperature
Temperature changes should be anticipated and expected under any general anesthetic and therefore any general anesthetic requires temperature measurement.
  Very brief procedures may be an exception, but the availability of temperature monitoring should be recorded.
Standards for Anesthetic Monitoring

## Slide 51: Anesthetic Monitoring
Temperature
  The temperature may be measured from many locations including:
    skin, nasopharynx, esophageal, bladder, rectal, or a pulmonary arterial catheter.
  Core temperatures obtained preferably from
    a pulmonary catheter,
    esophageal probe,
    or rectal probe
Standards for Anesthetic Monitoring

## Slide 52: Anesthetic Monitoring
Pulse Oximetry (Sp02)
Is one of the most commonly employed monitoring modalities in anesthesia.
It is a non-invasive way to monitor the oxygenation of a patient's hemoglobin.
A sensor with both red and infrared wavelengths is placed on the patient.
Absorption of these wavelengths by the blood is measured and oxygen saturation (Sp02) can be calculated.
Modalities for Anesthetic Monitoring

## Slide 53: Anesthetic Monitoring
Oximetry
Basic Concepts
There are two main types of oximetry : Fractional oximetry and Functional oximetry
Fractional oximetry. Oxyhemoglobin/(oxyhemoglobin + deoxyhemoglobin + methemoglobin + carboxyhemoglobin) Fractional oximetry measures the arterial oxygen saturation (SaO2)
Can only be measured by an arterial blood sample
Modalities for Anesthetic Monitoring

## Slide 54: Anesthetic Monitoring
Oximetry
Basic Concepts
There are two main types of oximetry : Fractional oximetry and Functional oximetry
Functional oximetry
Oxyhemoglobinl/(oxyhemoglobin + deoxyhemoglobin) = SpO2 Functional oximetry gives you the SpO2
Can be measured noninvasively by a standard pulse oximeter
Modalities for Anesthetic Monitoring

## Slide 55: Anesthetic Monitoring
How pulse oximetry works
A pulse oximeter emits two wavelengths of light: red (660 nm) and infrared (940 nm)
Deoxyhemoglobin absorbs more light in the red band
Oxyhemoglobin absorbs more light in the infrared band
b) Sensors in the oximeter detect the amount of red and infrared light absorbed by the blood
c) Photoplethysmography is then used to identify pulsatile arterial flow (alternating current [AC]) and non-pulsatile flow (direct current [DC])
d) The ratio of AC/DC at both 66o and 94o nm is measured using the equation: (AC/DC)660/(AC/DC)940
Modalities for Anesthetic Monitoring

## Slide 56: Anesthetic Monitoring
How pulse oximetry works
The pulse oximeter calculates the SpO2 by taking the above equation and using an algorithm built into the software to derive the SpO2
The calibration to derive SpO2 from the (AC/DC)660/(AC/DC)940 ratio was made from studies of healthy volunteers
Modalities for Anesthetic Monitoring

## Slide 57: Anesthetic Monitoring
Light absorption with Oxygenated and Deoxygenated Hemoglobin
Modalities for Anesthetic Monitoring

## Slide 58: Anesthetic Monitoring
Accuracy of the pulse oximeter
a) If the SpO2 is between 70% and 100%, the pulse oximeter is accurate to within 5%
b) It is not accurate below 70%
because calibration of the pulse oximeter involved healthy volunteers whose SpO2 did not routinely reach levels <70%
Modalities for Anesthetic Monitoring

## Slide 59: Anesthetic Monitoring
For the relationship between SaO2 and PaO2
The absorption spectrum of deoxygenated hemoglobin is very steep at 600 nm in the red range so small changes in the amount of deoxyhemoglobin can cause very wide variances in SpO2
Modalities for Anesthetic Monitoring

## Slide 60: Anesthetic Monitoring
For the relationship between SaO2 and PaO2
The absorption spectrum of deoxygenated hemoglobin is very steep at 600 nm in the red range so small changes in the amount of deoxyhemoglobin can cause very wide variances in SpO2
Modalities for Anesthetic Monitoring

## Slide 61: Anesthetic Monitoring
Low perfusion makes it difficult for the pulse oximeter to distinguish a true signal from background noise
Pulse oximetry is not as accurate in low amplitude states
Hypovolemia
Hypothermia
Cardiac arrest
Arrhythmias
Vasoconstriction
BP cuff inflation
Tourniquet
Cardiac Bypass
Modalities for Anesthetic Monitoring

## Slide 62: Anesthetic Monitoring
Dyshemoglobinemias
Pulse oximetry only accurately measures oxyhemoglobin and deoxyhemoglobin—all other forms of hemoglobin are not accurately measured
  Carboxyhemoglobin is measured as 90% oxyhemoglobin and 10% deoxyhemoglobin
    Thus, when there are high amounts of carboxyhemoglobin it will overestimate the SpO2
    This is an important consideration in patients exposed to smoke or fires
  Methemoglobin absorbs equal amounts of red and infrared light so the SpO2 will read 85%
    Methemoglobin is formed when iron goes from it's +2 ferrous form to the +3 ferric state
The ferric state of iron displays a left shift on the oxygen dissociation curve and releases oxygen less easily
Modalities for Anesthetic Monitoring

## Slide 63: Anesthetic Monitoring
Dyshemoglobinemias
Methemoglobinemia can be caused by many drugs.
Nitrates
Nitrites
Local anesthetics (eg Benzocaine)
Chlorates
Antimalarials
Antineoplastics
Sulfonamide
Dapsone (antibiotic)
Metoclopramide (antiemetic)
Modalities for Anesthetic Monitoring

## Slide 64: Anesthetic Monitoring
Dyshemoglobinemias
Patients with sickle cell anemia presenting in a vasoocclusive crisis can have an inaccurate SpO2 reading
High levels of bilirubin do not alter SpO2 readings
Modalities for Anesthetic Monitoring

## Slide 65: Anesthetic Monitoring
Normal capnogram
a) Phase I
  Initiation of expiration
  CO2 free gas from anatomic dead space
b) Phase II
Expiration of mixture of dead space and alveolar gas
c) Phase III
  Alveolar plateau
  CO2-rich gas from alveoli
d) Phase IV or 0
Inspiration
Standards for Anesthetic Monitoring

## Slide 66: Anesthetic Monitoring
Clinical uses of capnography
  Provides qualitative and quantitative information regarding expired CO2.
  Quantitatively, this is useful to ensure the endotracheal tube is within the respiratory tract
  Quantitatively , to ensure adequate cardiac output.
  Inspired CO2
  Monitored to ensure that the CO2 absorber of the anesthesia machine is adequately removing all CO2 from the circuit.
  If inspired CO2 is greater than zero, changing of the absorbent should be considered. The color of absorbent turns blue when its capacity is exhausted.
Standards for Anesthetic Monitoring

## Slide 67: Anesthetic Monitoring
Clinical uses of capnography
Monitoring of adequacy of ventilation in controlled or spontaneously ventilating patients
Noninvasive estimate of PaCO2
  Assumes the normal 2 to 5 mm Hg difference between expired (PETCO2) and arterial (PaCapnography in the awake state is present)
The gradient between PETCO2 and PaCO2 may be increased with age, pulmonary disease, pulmonary embolus, low cardiac output, and hypovolemia
Standards for Anesthetic Monitoring

## Slide 68: Anesthetic Monitoring
Clinical uses of capnography
Detection of patient disease
i) Causes of increased CO2 production
  Fever
  Sepsis
  Malignant hyperthermia
  Hyperthyroidism
  Shivering
ii) Causes of decreased PETCO2
  Decreased cardiac output
  Hypovolemia
  Pulmonary embolism
  Hypothermia
  Hyperventilation
Standards for Anesthetic Monitoring

## Slide 69: Anesthetic Monitoring
Detection of patient disease
Airway obstruction may be detected due to abnormalities in the capnography tracing.
Standards for Anesthetic Monitoring
Clinical uses of capnography

## Slide 70: Anesthetic Monitoring
Clinical uses of capnography
Detection of problems with the anesthetic breathing system
Rebreathingof CO2
Incompetent valves
Circuit disconnect
Circuit leak
Standards for Anesthetic Monitoring

## Slide 71: Anesthetic Monitoring
Clinical uses of capnography
Interpretation of abnormal capnograms
a) Rebreathing of CO2
  Elevation in baseline CO2 and Phase I
  Can eliminate by increasing fresh gas flow or changing CO2 absorber
b) Obstruction to expiratory gas flow
  Prolonged Phase II and steeper Phase III slope
  Occurs with bronchospasm, COPD, kinked endotracheal tube
Standards for Anesthetic Monitoring

## Slide 72: Anesthetic Monitoring
Clinical uses of capnography
Standards for Anesthetic Monitoring

## Slide 73: Anesthetic Monitoring
Clinical uses of capnography
Interpretation of abnormal capnograms
c) Curare Cleft
  Dip in Phase III
  Indicates return of spontaneous respiratory efforts
d) Cardiogenic oscillations
  Oscillations of small gas movements during phase III and IV (or 0)
  Produced by aortic and cardiac pulsations
Standards for Anesthetic Monitoring

## Slide 74: Anesthetic Monitoring
Clinical uses of capnography
Interpretation of abnormal capnograms
Increased CO2
  Elevated plateau height
  Indicates increased CO2 production states other source of CO2 (as in laparoscopic surgery), or inadequate minute ventilation
Decreased measured CO2
  Decreased plateau height
  May indicated decreased CO2 production state or increased minute ventilation
Standards for Anesthetic Monitoring

## Slide 75: Anesthetic Monitoring
Clinical uses of capnography
Interpretation of abnormal capnograms
Incompetent inspiratory valve
  Prolonged Phase III with elevation of baseline COz and plateau height
  Results in rebreathing
  May be difficult to detect without simultaneous analysis of flow waveforms
Esophageal intubation
  Initial presence of CO2 followed by no CO2
Standards for Anesthetic Monitoring

## Slide 76: Anesthetic Monitoring
Standards for Anesthetic Monitoring

## Slide 77: Anesthetic Monitoring
Standards for Anesthetic Monitoring

## Slide 78: Anesthetic Monitoring
Processed EEG and Awareness Monitoring
Intra-operative awareness with recall involves explicit recall of sensory perceptions during general anesthesia including aspects of their surgical environment, procedure, and even pain related to the intervention.
Intra-operative awareness with recall is defined as a patient having an unexpected and undesirable recall of wakefulness
Processed EEG analysis has been developed as a method to monitor depth of anesthesia intraoperatively and can be used as an effect-site monitor to aid in titration of anesthetic drugs and may be useful in reducing the incidence of intra-operative awareness with recall.
Modalities for Anesthetic Monitoring

## Slide 79: Anesthetic Monitoring
Processed EEG and Awareness Monitoring
Intraoperative awareness
Symptoms
The most common symptoms reported by patients suggesting awareness with recall are auditory perceptions such as voices or noises, followed by loss of motor function (inability to move, sensation of weakness, or paralysis), pain, and feelings of helplessness, anxiety, panic, impending death, or catastrophe.
Awareness with recall can lead to anxiety, sleep difficulties, insomnia, irritability, nightmares, and posttraumatic stress disorder.
Modalities for Anesthetic Monitoring

## Slide 80: Anesthetic Monitoring
Processed EEG and Awareness Monitoring
Incidence of awareness
The incidence of awareness with recall varies among studies, countries, anesthetic techniques, patient characteristics, and types of surgery.
The most commonly cited rate of intra-operative awareness is 0.2% .
This figure is thought to reflect the incidence in routine cases but not including cardiac or obstetric surgeries.
When further stratified, awareness occurs in approximately 1.14% to 1.5% of cardiac surgery cases, 0.4% of obstetric cases, and 11% to 43% of trauma surgeries.
Awareness with recall associated with pain is estimated to occur in 0.01% to 0.03% of cases.
Modalities for Anesthetic Monitoring

## Slide 81: Anesthetic Monitoring
Processed EEG and Awareness Monitoring
Factors associated with increased risk of awareness with recall include
"light" anesthesia (e.g., delivering a low level of inhaled anesthetic minimum alveolar concentration),
history of intra-operative awareness
chronic use of central nervous system depressants
younger age
obesity
inadequate or misused anesthesia delivery systems
Modalities for Anesthetic Monitoring

## Slide 82: Anesthetic Monitoring
Processed EEG and Awareness Monitoring
Detecting episodes of intra-operative awareness
Often it is difficult to know for sure that intra-operative awareness with recall occurred.
If the patient is not asked specifically about it they may not report it voluntarily.
Or, the patient may recollect hearing sounds during surgery, when in fact they are remembering something that occurred in the recovery room.
Modalities for Anesthetic Monitoring

## Slide 83: Anesthetic Monitoring
Processed EEG and Awareness Monitoring
Detecting episodes of intra-operative awareness
One accepted method to assess intra-operative awareness with recall is to conduct three structured interviews with open ended questions at intervals of 24 hours, between 24 and 72 hours, and at 30 days after surgery (awareness may not arise until days to weeks postoperatively).
Modalities for Anesthetic Monitoring

## Slide 84: Anesthetic Monitoring
Processed EEG and Awareness Monitoring
Prevention or vigilance for detecting intraoperative awareness
Monitor delivered volatile anesthetic levels
  The unintended inadequate delivery of volatile anesthetic agents ("light anesthesia") during maintenance of anesthesia may be avoided by the addition of a low alarm limit to end-tidal gas monitoring settings, as well as use of a "near empty" alarm in anesthetic vaporizers.
Modalities for Anesthetic Monitoring

## Slide 85: Anesthetic Monitoring
Processed EEG and Awareness Monitoring
Prevention or vigilance for detecting intraoperative awareness
Monitor processed EEG signals
  Depth of anesthesia monitoring, via the processed EEG, has proved useful in reducing the amount of anesthetic drugs, optimizing extubation times, and in some studies reducing awareness with recall.
  Although most anesthesiologists in the UK, USA, and Australia accept that clinical signs are unreliable indicators of awareness, few believe that monitors of anesthetic depths should be used for all routine cases
Modalities for Anesthetic Monitoring

## Slide 86: Anesthetic Monitoring
Processed EEG and Awareness Monitoring
Prevention or vigilance for detecting intraoperative awareness
Monitor processed EEG signals
  Several brain-function monitors based on the processed electroencephalogram (EEG) or evoked potentials have been developed to assess anesthetic depth.
Modalities for Anesthetic Monitoring

## Slide 87: Anesthetic Monitoring
Processed EEG and Awareness Monitoring
    BIS (Aspect Medical Systems). The most widely used monitor is the BIS monitor. This device integrates several parameters of an EEG into a calculated, dimensionless variable (0 to 100).
    It is important to note that bispectral index (BIS)is a probability distribution where a measure of 40 does not provide a 100% guarantee of no awareness.
Modalities for Anesthetic Monitoring

## Slide 88: Anesthetic Monitoring
Processed EEG and Awareness Monitoring
Several brain-function monitors based on the processed electroencephalogram
(EEG) or evoked potentials have been developed to assess anesthetic depth.
  The term bispectral applies because it incorporates both power and phase spectrums of an EEG into the calculated 0 to 100 value.
  BIS values between 40 and 60 purportedly indicate adequate general anesthesia for surgery, and values below 40 indicate a deep hypnotic state. Targeting a range of BIS values between 40 and 60 is marketed to help prevent anesthesia awareness while allowing for minimization the anesthetic dose.
Modalities for Anesthetic Monitoring

## Slide 89: Anesthetic Monitoring
Modalities for Anesthetic Monitoring

## Slide 90: Anesthetic Monitoring
Processed EEG and Awareness Monitoring
M-Entropy Module (GE-Healthcare). A mathematical approach that quantifies EEG using non-linear dynamics. This mode measures spectral entropy and applies it to the power spectrum of EEGs.
  Two variables, state and response entropy, which measure EEG and combined EEG/EMG activity respectively, are displayed on the awareness monitor as a dimensionless unit (0 to 100)
Mid-latency auditory evoked potentials (MLAEPs). This method is thought to be an alternative to the use of EEG monitoring.
MLAEP are electroencephalographic responses to auditory stimuli.
Modalities for Anesthetic Monitoring

## Slide 91: Anesthetic Monitoring
Modalities for Anesthetic Monitoring

## Slide 92: Anesthetic Monitoring

## Slide 93: Anesthetic Monitoring
Processed EEG and Awareness Monitoring
Modalities for Anesthetic Monitoring

## Slide 94: Anesthetic Monitoring
Processed EEG and Awareness Monitoring
Modalities for Anesthetic Monitoring

## Slide 95: Anesthetic Monitoring
Neurophysiologic Monitoring and
Anesthetic Management
Neurophysiologic monitoring or neuromonitoring allows early detection of events that may increase postoperative neurological morbidity.
The aim of monitoring is to identify changes in brain, spinal cord, and peripheral nerve function prior to irreversible damage.
Neuromonitoring is also useful in identifying anatomical structures.
Modalities for Anesthetic Monitoring

## Slide 96: Anesthetic Monitoring
Modalities for Anesthetic Monitoring

## Slide 97: Anesthetic Monitoring
Neurophysiologic Monitoring and
Anesthetic Management
Electromyography (EMG)
EMG is the recording of electrical activity of muscle and therefore an indirect indicator of function of the innervating peripheral nerve.
This technique is also used to identify and verify the integrity of a peripheral nerve, including cranial nerves as well as pedicle screw testing during spine surgery.
EMG is only sensitive to neuromuscular blocking agents.
Modalities for Anesthetic Monitoring

## Slide 98: Anesthetic Monitoring
Neurophysiologic Monitoring and
Anesthetic Management
Somatosensory evoked potentials (SSEP)
SSEP are the recording, usually at the cerebral cortex, of responses from electrically stimulated peripheral afferent nerves.
The most commonly used peripheral nerves are median, ulnar, posterior tibial, and common peroneal nerves.
Modalities for Anesthetic Monitoring

## Slide 99: Anesthetic Monitoring
Neurophysiologic Monitoring and
Anesthetic Management
Brainstem auditory evoked potentials (BAEP)
  BAEP are the recording of brainstem responses to auditory stimuli.
  BAEP monitors the function of the entire auditory pathway along the acoustic nerve, through the brain stem to the cerebral cortex.
Modalities for Anesthetic Monitoring

## Slide 100: Anesthetic Monitoring
Neurophysiologic Monitoring and
Anesthetic Management
Motor evoked potentials (MEP)
MEP is the recording obtained from electrical stimulation of the motor cortex, which elicits potentials in the spinal cord or (myogenic) potentials from the innervated muscle.
Monitors motor pathway function
Modalities for Anesthetic Monitoring

## Slide 101: Anesthetic Monitoring
Neurophysiologic Monitoring and
Anesthetic Management
Electroencephalography (EEG)
EEG monitoring can be a useful supplement to surgery when
  Seizure foci need to be identified
  The general state of cerebral metabolism needs monitoring
  Cerebral ischemia can occur
EEG is a standard of care in many institutions for carotid endarterectomy.
Modalities for Anesthetic Monitoring

## Slide 102: Anesthetic Monitoring
Neurophysiologic Monitoring and
Anesthetic Management
Electroencephalography (EEG)
  iii) EEG is the recording of brain electrical activity and is highly dependent on anesthetic depth.
      (I) Alpha waves are rhythmically regular waves of 8 to 12 Hz seen in a lightly anesthesized patient.
        A faster, disorganized beta (>12 Hz) rhythm is seen upon awakening.
        Slower theta waves (4 to 8 Hz) are seen with deep inhalation or moderate dose narcotic anesthesia.
Slow delta waves (<4 Hz) indicate deep anesthesia, or ischemia if the amplitude is low.
Modalities for Anesthetic Monitoring

## Slide 103: Anesthetic Monitoring
Modalities for Anesthetic Monitoring

## Slide 104: Anesthetic Monitoring
  CEREBRAL OXIMETRY
  Cerebral oximetry uses near infrared spectroscopy (NIRS).
  Using reflectance spectroscopy near infrared light is emitted by a probe on the scalp
  Receptors are likewise positioned to detect the reflected light from both deep and superficial structures.
  As with pulse oximetry, oxygenated and deoxygenated hemoglobin absorb light at different frequencies.
  Likewise, cytochrome absorbs infrared light in the mitochondria.
  The NIRS saturation largely reflects the absorption of venous hemoglobin, as it does not have the ability to identify the pulsatile arterial component.
  Regional saturations of less than 40% on NIRS measures, or changes of greater than 25% of baseline measures, may herald neurological events secondary to decreased cerebral oxygenation.
Modalities for Anesthetic Monitoring

## Slide 105: Anesthetic Monitoring
  CEREBRAL OXIMETRY
Modalities for Anesthetic Monitoring

## Slide 106: Anesthetic Monitoring
  Invasive pressure monitoring
  Central Venous Pressure
  Central venous catheterization involves placement of a sterile catheter into one of the large central veins and allows for multiple modalities of intervention along with the option of monitoring central venous pressure (CVP).
  CVP monitoring can be a useful tool for evaluating intravascular volume and preload in the absence of left ventricular (LV) dysfunction (ejection fraction <40%), severe mitral valve disease, pulmonary hypertension, or significant reduction in LV compliance (ischemia/diastolic dysfunction).
Modalities for Anesthetic Monitoring

## Slide 107: Anesthetic Monitoring
  Invasive pressure monitoring
  Central Venous Pressure
Modalities for Anesthetic Monitoring

## Slide 108: Anesthetic Monitoring
  Invasive pressure monitoring
  Central Venous Pressure
Modalities for Anesthetic Monitoring

## Slide 109: Anesthetic Monitoring
  Invasive pressure monitoring
  Pulmonary artery Pressure
  The pulmonary artery (PA) catheter is a controversial but potentially powerful tool, offering information about cardiac filling pressures, cardiac output (CO), derived parameters of cardiac performance, and mixed venous oxygen saturation (Sv02).
  ASA consensus opinion is that "PA catheter monitoring may reduce perioperative complications if critical hemodynamic data obtained are accurately interpreted and appropriate treatment is instituted.
Modalities for Anesthetic Monitoring

## Slide 110: Anesthetic Monitoring
  Invasive pressure monitoring
  Pulmonary artery catheter
Modalities for Anesthetic Monitoring

## Slide 111: Anesthetic Monitoring
  Invasive pressure monitoring
  Pulmonary artery Pressure
Modalities for Anesthetic Monitoring

## Slide 112: Anesthetic Monitoring
  Invasive pressure monitoring
  Pulmonary artery Pressure
Modalities for Anesthetic Monitoring

## Slide 113: Anesthetic Monitoring
  TransEsophageal Echocardiography
Transesophageal echocardiograpy (TEE) is a monitoring modality gaining popularity in the field of anesthesiology due to its versatility, reliability, and safety.
It was initially used as a diagnostic tool primarily by cardiologists but has become a mainstay in intraoperative cardiac anesthesia and its utility is extending into other areas as well.
Modalities for Anesthetic Monitoring

## Slide 114: Anesthetic Monitoring
PERIPHERAL NERVE STIMULATION
Indications
Because of the variation in patient sensitivity to neuromuscular blocking agents, the neuromuscular function of all patients receiving intermediate- or long-acting neuromuscular blocking agents should be monitored.
In addition, peripheral nerve stimulation is helpful in assessing paralysis during rapid-sequence inductions or during continuous infusions of short-acting agents.
Furthermore, peripheral nerve stimulators can help locate nerves to be blocked by regional anesthesia.
Modalities for Anesthetic Monitoring

## Slide 115: Anesthetic Monitoring
PERIPHERAL NERVE STIMULATION
Contraindications
There are no contraindications to neuromuscular monitoring, although certain sites may be precluded by the surgical procedure.
Additionally, atrophied muscles in areas of hemiplegia or nerve damage may appear refractory to neuromuscular blockade secondary to the proliferation of receptors.
Determining the degree of neuromuscular blockade using such an extremity could lead to potential overdosing of competitive neuromuscular blocking agents.
Modalities for Anesthetic Monitoring

## Slide 116: Anesthetic Monitoring
PERIPHERAL NERVE STIMULATION
Techniques & Complications
A peripheral nerve stimulator delivers current (60- 80 mA) to a pair of either ECG silver chloride pads or subcutaneous needles placed over a peripheral motor nerve.
The evoked mechanical or electrical response of the innervated muscle is observed.
Although electromyography provides a fast, accurate, and quantitative measure of neuromuscular transmission, visual or tactile observation of muscle contraction is usually relied upon in clinical practice.
Ulnar nerve stimulation of the adductor pollicis muscle and facial nerve stimulation of the orbicularis oculi are most commonly monitored.
Modalities for Anesthetic Monitoring

## Slide 117: Anesthetic Monitoring
PERIPHERAL NERVE STIMULATION
Modalities for Anesthetic Monitoring

## Slide 118: Anesthetic Monitoring
PERIPHERAL NERVE STIMULATION
Techniques & Complications
Because it is the inhibition of the neuromuscular receptor that needs to be monitored, direct stimulation of muscle should be avoided by placing electrodes over the course of the nerve and not over the muscle itself.
Complications of nerve stimulation are limited to skin irritation and abrasion at the site of electrode attachment.
Modalities for Anesthetic Monitoring

## Slide 119: Anesthetic Monitoring
PERIPHERAL NERVE STIMULATION
Techniques & Complications
Because of concerns of residual neuromuscular blockade, increased attention has been focused on providing quantitative measures of the degree of neuromuscular blockade perioperatively.
Acceleromyography uses a piezoelectric transducer on the muscle to be stimulated. Movement of the muscle generates an electrical current that can be quantified and displayed.
Indeed, acceleromyography can better predict residual paralysis, compared with routine tactile train-of-four monitoring used in most operating rooms, if calibrated from the beginning of the operative period to establish baselines prior to administration of neuromuscular blocking agents.
Modalities for Anesthetic Monitoring

## Slide 120: Anesthetic Monitoring
PERIPHERAL NERVE STIMULATION
Clinical Considerations
The degree of neuromuscular blockade is monitored by applying various patterns of electrical stimulation
All stimuli are 200 µs in duration and of square-wave pattern and equal current intensity.
A twitch is a single pulse that is delivered from every 1 to every 10 sec (1–0.1 Hz).
Increasing block results in decreased evoked response to stimulation.
Modalities for Anesthetic Monitoring

## Slide 121: Anesthetic Monitoring
PERIPHERAL NERVE STIMULATION
Clinical Considerations
Train-of-four stimulation denotes four successive 200-µs stimuli in 2 sec (2 Hz).
The twitches in a train-of-four pattern progressively fade as nondepolarizing muscle relaxant block increases. The ratio of the responses to the first and fourth twitches is a sensitive indicator of nondepolarizing muscle paralysis. Ratio o fourth twitch over the first twitch should be greater than or equal to 90% to give the reversal ( neostigmine and glycopyrrolate)
Because it is difficult to estimate the train-of-four ratio, it is more convenient to visually observe the sequential disappearance of the twitches, as this also correlates with the extent of blockade. Disappearance of the fourth twitch represents a 75% block, the third twitch an 80% block,
and the second twitch a 90% block.
Clinical relaxation usually requires 75% to 95% neuromuscular blockade.
Modalities for Anesthetic Monitoring

## Slide 122: Anesthetic Monitoring
PERIPHERAL NERVE STIMULATION
Clinical Considerations
Train-of-four stimulation denotes four successive 200-µs stimuli in 2 sec (2 Hz).
blockade.
Modalities for Anesthetic Monitoring

## Slide 123: Anesthetic Monitoring
PERIPHERAL NERVE STIMULATION
Clinical Considerations
Tetany at 50 or 100 Hz is a sensitive test of neuromuscular function. Sustained contraction for 5 sec indicates adequate—but not necessarily complete— reversal from neuromuscular blockade.
Double-burst stimulation (DBS) represents two variations of tetany that are less painful to the patient.
The DBS 3,3 pattern of nerve stimulation consists of three short (200-µs) high-frequency bursts separated by 20 ms intervals (50 Hz) followed 750 ms later by another three bursts.
DBS 3,2 consists of three 200-µs impulses at 50 Hz followed 750 ms later by two such impulses.
DBS is more sensitive than train-of-four stimulation for the clinical (ie, visual) evaluation of fade.
Modalities for Anesthetic Monitoring

## Slide 124: Anesthetic Monitoring
PERIPHERAL NERVE STIMULATION
Clinical Considerations
Single Twitch
Train of four
Double Burst Stimulation
Post Tetanic Count
Modalities for Anesthetic Monitoring

## Slide 125: Anesthetic Monitoring
PERIPHERAL NERVE STIMULATION
Clinical Considerations
Because muscle groups differ in their sensitivity to neuromuscular blocking agents, use of the peripheral nerve stimulator cannot replace direct observation of the muscles (eg, the diaphragm) that need to be relaxed for a specific surgical procedure.
Furthermore, recovery of adductor pollicis function does not exactly parallel recovery of muscles required to maintain an airway.
Modalities for Anesthetic Monitoring

## Slide 126: Anesthetic Monitoring
PERIPHERAL NERVE STIMULATION
Clinical Considerations and Extubation Criteria
The diaphragm, rectus abdominis, laryngeal adductors, and orbicularis oculi muscles recover from neuromuscular blockade sooner than do the adductor pollicis.
Other indicators of adequate recovery include sustained (≥5 s) head lift, the ability to generate an inspiratory pressure of at least –25 cm H 2 O, and a forceful hand grip.
Twitch tension is reduced by hypothermia of the monitored muscle group (6%/°C).
Decisions regarding adequacy of reversal of neuromuscular blockade, as well as timing of extubation, should be made only by considering both the patient’s clinical presentation and assessments determined by peripheral nerve stimulation.
Modalities for Anesthetic Monitoring

## Slide 127: Anesthetic Monitoring
PERIPHERAL NERVE STIMULATION
Clinical Considerations
Postoperative residual curarization remains a problem in post-anesthesia care, producing potentially injurious airway and respiratory function compromise.
Reversal of neuromuscular blocking agents is warranted, as is the use of intermediate acting neuromuscular blocking agents instead of longer acting drugs.
Modalities for Anesthetic Monitoring

## Slide 128
Modalities for Anesthetic Monitoring

## Slide 129: Anesthetic Monitoring
  Electrolytes/Acid Base
  Coagulation
  Urine output
Modalities for Anesthetic Monitoring